## Slide 1
P39
P25
Application Guide
Coatron A6 Plus
Version: Rev1d
Firmware: V1.03.05d (50)
Issued: May 2019
Document ID: xxxxxxx

## Slide 2
General Information
This book is intended as an essential guide of how and where to place the reagents onboard and how to setup the instrument. Further and detailed information can be read in the operation manual of instrument or the box insert of reagent. The instrument protocol parameters stated in this book have been developed and validated by TECO GmbH to ensure highest performance and accuracy. Any modifications of these parameters are completely under the responsibility of the user and can impact the quality of measurement. A full validation is required after any change !
History of versions |  |  | 
Version | Firmware | Date | Comment
Rev-1c | 1.03.05b | 12.07.2018 | Initial version
Rev-1d | 1.03.05d (50) | 16.05.2019 | 
 | Copyright © 2018 TECO Medical Instruments, Production + Trading GmbH Dieselstrasse. 1; D-84088 Neufahrn in N.B in Germany

## Slide 3
Intended use
Use only human citrated plasma for sample analysis. Mix 9 parts of venous blood with 1 part 3.2% (0.105M) sodium citrate and centrifuge the mixture at 1500g x 15min. Use plasma within 4 hours. Do not use plasma with more than 50mg/dL Bilirubin (856µmol/l) or 2000mg/L Hemoglobin or 50 g/l Triglyceride (57 mmol/l) The enviromental temperature should be conditioned to 20 – 25°C to allow accurate cooling and best reagent onboard stability Do not store reagents over night onboard in uncapped condition.
Safety Information
Never touch moving parts such as the measurement rotor or pipetting arm during device operation. Never try to pull a cuvette block out of the measurement rotor during test processing operation. Carry out control measurement runs at regular intervals to ensure that the Analyzer continues to function faultlessly Avoid direct contact with samples and sample residues in the used cuvettes. Infectious material such as cuvette waste and liquid waste must be disposed in compliance with local regulations governing for infectious materials. Wear medical infection grade protective gloves for all cleaning and maintenance works involving potential contact with infectious liquids and use each pair of gloves once only. Use a hand disinfectant product, e.g. Sterilium®, to disinfect your hands after completion of the work.

## Slide 4
Glossary of test protocol parameters |  | 
Test | PT | Name of test
UNIT | 251 | Available result units (s+%+INR+..) ; decimal code of every bit
S-CORR | 0% | Signal correction
T-CORR | 30%, -4 s | Result correction ( eg. correct clotting time +30% and substract 4s)
START | 7 s | Start of measurement
INCUB | 0 s | Incubation time between pipet step R1 und R2
RUNTIME | 120 s | End of measurement
METHOD | COAG | Select between clotting, chromogenic or immunturbidimetric test method
MATH | 1/X | Interpolation of calibration curve
CT-MECH | NO | Definition of clot; NO = clot at turnpoint YES = clot at halfpoint of reaction
SENS / LIMIT | HI | Sensitivity of clotting method / Limit of chromogenic method
MIX | 0 | After adding of start reagent the robot will mix in the cuvette
CLEAN | 0 | Enable high cleaning cycle ( e.g. required for thrombin reagents)
POINTS | 4 | Number of auto calibration points (100% , 50%, 25%, 12,5%, 6,25%, 1,25%)
MULTI | 3 | Enable multi dispensing ( one input, several outputs).
PAT | 25µL | Patient plasma (volume)
BUF | 0 | Buffer (sample diluent)
CLR | 0 | Clear Volume ( eg. remove 150µL sample dilution from cuvette)
DEF | 0 | Deficient plasma
R0 | 0 | Reagent-0
R1 | 0 | Reagent-1
R2 | 50µL P25 | Reagent-2 ( = start reagent)

## Slide 5
Instrument | Coatron A6 Plus | 
Firmware | V1.03.05d (50) | 
 | Volume | Position
PAT | 25µl | 
BUF | 0 | P79
CLR | 0 | 
DEF | 0 | 
R0 | 0 | 
R1 | 0 | 
R2 | 50ul | P46
REF | Determinations | 
A0230-010 | 5x2ml | n=200
A0230-040 | 10x4ml | n=800
A0230-100 | 10x10ml | n=2000
Test | PT
UNIT | 251
S-CORR | 0%
T-CORR | 30%, -4 s
START | 7 s
INCUB | 0 s
RUNTIME | 90 s
METHOD | COAG
MATH | 1/X
CT-MECH | NO
SENS | HI
MIX | 0
CLEAN | 0
POINTS | 4
MULTI | 3
TECLOT PT-S
REF: A0230-xxx
Test Protocol
P46: PT reagent
P79: IBS buffer
(only for auto-calibration)

## Slide 6
Auto Calibration of PT-% |  |  | 
Material |  |  | 
Calibrator | TECAL N |  | 1ml
Diluent | IBS Buffer |  | 2ml
Tubes | 13x75 mm |  | 4x
Preparation |  |  | 
1 | Reconstitute PT reagent and place it onboard |  | 
2 | Reconstitute calibrator and place into P01 |  | 
3 | Place 3 empty tubes into P02 – P04 |  | 
4 | Wait for 60 mins |  | 
Procedure |  |  | 
1 | Open menu SETUP TEST and select PT |  | 
2 | Enter LOT and expiry date by hand or barcode |  | 
3 | Select UNIT = % |  | 
4 | Select ENTRY=AUTO |  | 
5 | Enter reference value of calibrator |  | 
6 | Run measurement |  | 
7 | Verify and save calibration curve |  | 
Typical calibration curve (The laboratory must establish its own curve for each new LOT of reagent ) |  |  | 
1 | 100 % | 15.0 sec | 
2 | 50 % | 20.0 sec | 
3 | 25 % | 33.0 sec | 
4 | 12,5 % | 58.0 sec | 
TECLOT PT-S
REF: A0230-xxx

## Slide 7
PT-INR |  | 
Note: Two ways for obtaining MNPT: Method 1: Run PT-% Calibration curve, instrument will automatically calculate the 100% (MNPT) |  | 
Method 2: Test at least 20 samples from healthy people and calculate mean value (MNPT) of PT results |  | 
Material for method 1: |  | 
Calibrator | TECAL N | 1ml
Diluent | IBS Buffer | 2ml
Tubes | 13x75 mm | 4x
Preparation |  | 
1 | Reconstitute PT reagent and place it onboard | 
2 | Reconstitute calibrator and place into P01 | 
3 | Place 3 empty tubes into P02 – P04 | 
4 | Wait for 60 mins | 
Procedure |  | 
1 | Open menu SETUP TEST and select PT | 
2 | Enter LOT and expiry date by hand or barcode | 
3 | Select UNIT = %, INR | 
4 | Select ENTRY=Auto | 
5 | Run PT Calibration | 
6 | Enter MNPT and ISI | 
7 | Verify and save data | 
TECLOT PT-S
REF: A0230-xxx

## Slide 8
Instrument | Coatron A6 Plus | 
Firmware | V1.03.05d (50) | 
 | Volume | Position
PAT | 25 | 
BUF | 0 | P79
CLR | 0 | 
DEF | 0 | 
R0 | 0 | 
R1 | 25 | P60
R2 | 25 | P47
REF | Determinations | 
A03xx-025 | 5x5ml | n=1000
A03xx-050 | 10x5ml | n=2000
A03xx-100 | 10x10ml | n=4000
Test | APTT
UNIT | 17
S-CORR | 0%
T-CORR | 15%, -6 s
START | 21 s
INCUB | 120 s
RUNTIME | 120 s
METHOD | COAG
MATH | lin
CT-MECH | NO
SENS | LO
MIX | 0
CLEAN | 0
POINTS | 3
MULTI | 3
TECLOT aPTT-S
REF: A03x-xxx
Test Protocol
P47: CaCl2
P60: aPTT reagent

## Slide 9
aPTT- Ratio | 
Preparation | 
1 | Place APTT and CaCl2 reagent onboard
2 | Wait for 30 mins
Procedure | 
1 | Open menu SETUP TEST and select APTT
2 | Enter LOT and expiry date by hand or barcode
3 | Select UNIT = R
4 | Select ENTRY=Manual
5 | Enter normal time
6 | Verify and save
TECLOT aPTT-S
REF: A03x-xxx

## Slide 10
Instrument | Coatron A6 Plus | 
Firmware | V1.03.05d (50) | 
 | Volume | Position
PAT | 30 | 
BUF | 0 | P79
CLR | 0 | 
DEF | 0 | 
R0 | 0 | 
R1 | 30 | P60
R2 | 30 | P47
REF | Determinations | 
A03xx-025 | 5x5ml | n=833
A03xx-050 | 10x5ml | n=1666
A03xx-100 | 10x10ml | n=3333
Test | HEP
UNIT | 1025
S-CORR | 0%
T-CORR | 0%, 0 s
START | 21 s
INCUB | 180 s
RUNTIME | 120 s
METHOD | COAG
MATH | lin
CT-MECH | NO
SENS | LO
MIX | 0
CLEAN | 0
POINTS | 3
MULTI | 3
Heparin (aPTT-S) REF: A03x-xxx
Test Protocol
P47: CaCl2
P60: aPTT reagent

## Slide 11
Auto Calibration of HEP IU/mL |  | 
Material |  | 
Calibrator | Enoxaparin like Clexane® 2.000 I. E. (20 mg)/0,2 ml Injektionslösung | 0,2ml
Diluent 1 | 0,9% NaCl | 200ml 2x 1mL
Diluent 2 | TEControl N | 
Tubes | 13x75 mm | 5x
Preparation |  | 
1 | place APTT and CaCl2 reagents onboard | 
2 | Dilute Clexane down to 10 IE/mL (1:1000) with 0.9% NaCL 10 IE/mL = 0.2mL Clexane + 200mL NaCl | 
3 | Prepare 5 heparin calibrators with different levels and place onboard P01 – P05 1.0 IE = 50µL 10IE + 450µL Tecontrol N 0.5 IE = 200µL 1.0 IE + 200µL Tecontrol N 0.25 IE = 200µL 0.5 IE + 200µL Tecontrol N 0.125 IE = 200µL 0.25 IE + 200µL Tecontrol N 0 IE = 400µL Tecontrol N | 
4 | Wait for 30 mins | 
Procedure |  | 
1 | Open menu SETUP TEST and select HEP | 
2 | Enter LOT and expiry date by hand or barcode | 
3 | Select UNIT = IU/mL | 
4 | Select ENTRY=AUTO | 
5 | Now Input 5 target values (edit, press ENTER, press DOWN) STD1 = 1.000 STD2 = 0.500 STD3 = 0.250 STD4 = 0.125 STD5 = 0.001 | 
6 | Run measurement | 
7 | Verify and save calibration curve | 
Heparin (based on aPTT)
REF: A0230-xxx

## Slide 12
Instrument | Coatron A6 Plus | 
Firmware | V1.03.05d (50) | 
 | Volume | Position
PAT | 50µl | 
BUF | 0 | P79
CLR | 0 | 
DEF | 0 | 
R0 | 0 | 
R1 | 0 | 
R2 | 25ul | P50
REF | Determinations | 
A0401-020 | 10x2ml | n=800
Test | TT
UNIT | 17
S-CORR | 0%
T-CORR | 20%, -3 s
START | 7 s
INCUB | 0 s
RUNTIME | 90 s
METHOD | COAG
MATH | lin
CT-MECH | NO
SENS | HI
MIX | 0
CLEAN | 0
POINTS | 0
MULTI | 3
TECLOT TT
REF: A0401-020
Test Protocol
P50: TT reagent

## Slide 13
TECLOT TT
REF: A0401-020
TT- Ratio | 
Preparation | 
1 | Reconstitute TT reagent and place it onboard
2 | Wait for 30 mins
Procedure | 
1 | Open menu SETUP TEST and select TT
2 | Enter LOT and expiry date by hand or barcode
3 | Select UNIT = R
4 | Select ENTRY=Manual
5 | Enter Normal value
6 | Verify and save

## Slide 14
Instrument | Coatron A6 Plus | 
Firmware | V1.03.05d (50) | 
 | Volume | Position
PAT | 7 | 
BUF | 53 | P79
CLR | 0 | 
DEF | 0 | 
R0 | 0 | 
R1 | 0 | 
R2 | 30 | P49
REF | Determinations | 
A0501-010 | 5x2ml | n=333
A0501-025 | 5x5ml | n=833
A0511-020 | 10x2ml | n=666
A0511-050 | 10x5ml | n=1666
Test | FIB
UNIT | 769
S-CORR | 0%
T-CORR | 0%, 0 s
START | 3 s
INCUB | 0 s
RUNTIME | 90 s
METHOD | COAG
MATH | logXY
CT-MECH | YES, 50%
SENS | MIN
MIX | 0
CLEAN | 2
POINTS | 4
MULTI | 3
TECLOT FIB
REF: A05xx-xxx
Test Protocol
P49: FIB reagent
P79: IBS buffer

## Slide 15
Auto Calibration of FIB-mg/dL (g/L) |  |  | 
Material |  |  | 
Calibrator | TECAL N |  | 1ml
Diluent | IBS Buffer |  | 2ml
Tubes | 13x75 mm |  | 3x
Preparation |  |  | 
1 | Reconstitute FIB reagent and place it onboard |  | 
2 | Reconstitute calibrator and place into P01 |  | 
3 | Place 2 empty tubes into P02 and P03 |  | 
4 | Wait for 30 mins |  | 
Procedure |  |  | 
1 | Open menu SETUP TEST and select FIB |  | 
2 | Enter LOT and expiry date by hand or barcode |  | 
3 | Select UNIT = mg/dL or g/L |  | 
4 | Select ENTRY=AUTO |  | 
5 | Enter reference value of calibrator |  | 
6 | Run measurement |  | 
7 | Verify and save calibration curve |  | 
Typical calibration curve (The laboratory must establish its own curve for each new LOT of reagent ) |  |  | 
1 | 500 mg/dL | 5.7 sec | 
2 | 250 mg/dL | 9.4 sec | 
3 | 125 mg/dL | 14.9 sec | 
4 | 83 mg/dL | 20.3 sec | 
TECLOT FIB
REF: A05xx-xxx

## Slide 16
Instrument | Coatron A6 Plus | 
Firmware | V1.03.05d (50) | 
 | Volume | Position
PAT | 4 | 
BUF | 36 | P79
CLR | 0 | 
DEF | 0 | 
R0 | 0 | 
R1 | 25 | P57
R2 | 100 | P64
REF | Determinations | 
C1010-020 | kit | n=360
Test | AT
UNIT | 34819
S-CORR | 0%
T-CORR | 0%, 0 s
START | 15 s
INCUB | 0 s
RUNTIME | 90 s
METHOD | CHROM
MATH | lin
CT-MECH | NO
LIMIT | 350 mE
MIX | 0
CLEAN | 1
POINTS | 3
MULTI | 0
TEChrom AT liquid (anti-Xa)
REF: C1010-020
Test Protocol
P64: Factor Xa reagent
P57: Factor Xa Substrate
P79: IBS buffer

## Slide 17
Auto Calibration of AT-% |  |  | 
Material |  |  | 
Calibrator | TECAL N |  | 1ml
Diluent | IBS Buffer |  | 2ml
Tubes | 13x75 mm |  | 3x
Preparation |  |  | 
1 | Place AT liquid reagent onboard |  | 
2 | Reconstitute calibrator and place into P01 |  | 
3 | Place 2 empty tubes into P02 and P03 |  | 
4 | Wait for 30 mins |  | 
Procedure |  |  | 
1 | Open menu SETUP TEST and select AT |  | 
2 | Enter LOT and expiry date by hand or barcode |  | 
3 | Select UNIT = % |  | 
4 | Select ENTRY=AUTO |  | 
5 | Enter reference value of calibrator |  | 
6 | Run measurement |  | 
7 | Verify and save calibration curve |  | 
Typical calibration curve (The laboratory must establish its own curve for each new LOT of reagent ) |  |  | 
1 | 102 % | 0,092 E | 
2 | 51 % | 0,279 E | 
3 | 25.5 % | 0,369 E | 
TEChrom AT liquid (anti-Xa)
REF: C1010-020

## Slide 18
Instrument | Coatron A6 Plus | 
Firmware | V1.03.05d (50) | 
 | Volume | Position
PAT | 3 | 
BUF | 116 | P78
CLR | 90 | 
DEF | 0 | 
R0 | 0 | 
R1 | 40 | P64
R2 | 40 | P57
REF | Determinations | 
C1000-010 | kit | n=250
Test | AT
UNIT | 34819
S-CORR | 0%
T-CORR | 0%, 0 s
START | 25 s
INCUB | 0 s
RUNTIME | 90 s
METHOD | CHROM
MATH | lin
CT-MECH | NO
LIMIT | 350mE
MIX | 0
CLEAN | 1
POINTS | 3
MULTI | 0
TEChrom AT (anti-Xa)
REF: C1000-010
Test Protocol
P64: Factor Xa reagent
P57: Factor Xa Substrate
P78: Sample Diluent
(1:5 working solution)

## Slide 19
Auto Calibration of AT-% |  | 
Material |  | 
Calibrator | TECAL N | 1ml
Diluent | Sample dilution (1:5 working solution) | 5ml
Tubes | 13x75 mm | 3x
Preparation |  | 
1 | Place AT reagent onboard | 
2 | Reconstitute calibrator and place into P01 | 
3 | Place 2 empty tubes into P02 and P03 | 
4 | Wait for 30 mins | 
Procedure |  | 
1 | Open menu SETUP TEST and select AT | 
2 | Enter LOT and expiry date by hand or barcode | 
3 | Select UNIT = % | 
4 | Select ENTRY=AUTO | 
5 | Enter reference value of calibrator | 
6 | Run measurement | 
7 | Verify and save calibration curve | 
TEChrom AT (anti-Xa)
REF: C1000-010 (5x2mL)

## Slide 20
Instrument | Coatron A6 Plus | 
Firmware | V1.03.05d (50) | 
 | Volume | Position
PAT | 10 | 
BUF | 40 | P52
CLR | 0 | 
DEF | 0 | 
R0 | 0 | 
R1 | 0 | 
R2 | 30 | P53
REF | Determinations | 
D2020-005 | kit | n=166
D2020-010 | kit | n=333
Test | DD
UNIT | 30721
S-CORR | -15%
T-CORR | 0%, 0 s
START | 60 s
INCUB | 0 s
RUNTIME | 180 s
METHOD | IMMUN
MATH | lin
CT-MECH | NO
LIMIT | 50 mE
MIX | 1
CLEAN | 2
POINTS | 2
MULTI | 0
BLUE D-Dimer LC
REF: D2020-xxx
Test Protocol
P52: Reaction Buffer
P53: Latex
P79: IBS buffer
(only for auto-calibration)

## Slide 21
Auto Calibration of DD-ng/mL or µg/mL, µg/L, mg/L |  |  | 
Material |  |  | 
Calibrator | TECAL DD |  | 1ml
Diluent | IBS Buffer |  | 2ml
Tubes | 13x75 mm |  | 2x
Preparation |  |  | 
1 | Place DD reagent onboard |  | 
2 | Reconstitute calibrator and place into P01 |  | 
3 | Place 1 empty tubes into P02 |  | 
4 | Wait for 30 mins |  | 
Procedure |  |  | 
1 | Open menu SETUP TEST and select DD |  | 
2 | Enter LOT and expiry date by hand or barcode |  | 
3 | Select UNIT = ng/mL or µg/mL, µg/L, mg/L |  | 
4 | Select ENTRY=AUTO |  | 
5 | Enter reference value of calibrator |  | 
6 | Run measurement |  | 
7 | Verify and save calibration curve |  | 
Typical calibration curve (The laboratory must establish its own curve for each new LOT of reagent ) |  |  | 
1 | 3413 ng/mL | 0,237 E | 
2 | 1706 ng/mL | 0,115 E | 
BLUE D-Dimer LC
REF: D2020-xxx

## Slide 22
Instrument | Coatron A6 Plus | 
Firmware | V1.03.05d (50) | 
 | Volume | Position
PAT | 8 | 
BUF | 32 | P79
CLR | 0 | 
DEF | 40 | P58
R0 | 0 | 
R1 | 0 | 
R2 | 80 | P46
REF | Determinations | 
P5001-010 | 10x1ml | n=250
Test | F2
UNIT | 32771
S-CORR | 0%
T-CORR | 0%, 0 s
START | 9 s
INCUB | 120 s
RUNTIME | 180 s
METHOD | COAG
MATH | logXY
CT-MECH | NO
SENS | HI
MIX | 0
CLEAN | 0
POINTS | 6
MULTI | 0
Factor II Deficient Plasma
REF: P5001-010
Test Protocol
P46: PT reagent
P58: Factor II deficient plasma
P79: IBS buffer

## Slide 23
Auto Calibration of FII-% |  | 
Material |  | 
Calibrator | TECAL N | 1ml
Diluent | IBS Buffer | 5ml
Tubes | 13x75 mm | 6x
Preparation |  | 
1 | Reconstitute Factor II deficient plasma and PT reagent, place them onboard | 
2 | Reconstitute calibrator and place into P01 | 
3 | Place 5 empty tubes into P02 – P06 | 
4 | Wait for 60 mins | 
Procedure |  | 
1 | Open menu SETUP TEST and select FII | 
2 | Enter LOT and expiry date by hand or barcode | 
3 | Select UNIT = % | 
4 | Select ENTRY=AUTO | 
5 | Enter reference value of calibrator | 
6 | Run measurement | 
7 | Verify and save calibration curve | 
Factor II Deficient Plasma
REF: P5001-010

## Slide 24
Instrument | Coatron A6 Plus | 
Firmware | V1.03.05d (50) | 
 | Volume | Position
PAT | 8 | 
BUF | 32 | P79
CLR | 0 | 
DEF | 40 | P59
R0 | 0 | 
R1 | 0 | 
R2 | 80 | P46
REF | Determinations | 
P5101-010 | 10x1ml | n=250
Test | F5
UNIT | 32771
S-CORR | 0%
T-CORR | 0%, 0 s
START | 9 s
INCUB | 120 s
RUNTIME | 180 s
METHOD | COAG
MATH | logXY
CT-MECH | NO
SENS | HI
MIX | 0
CLEAN | 0
POINTS | 6
MULTI | 0
Factor V Deficient Plasma
REF: P5101-010
Test Protocol
P46: PT reagent
P59: Factor V deficient plasma
P79: IBS buffer

## Slide 25
Auto Calibration of FV-% |  | 
Material |  | 
Calibrator | TECAL N | 1ml
Diluent | IBS Buffer | 5ml
Tubes | 13x75 mm | 6x
Preparation |  | 
1 | Reconstitute Factor V deficient plasma and PT reagent, place them onboard | 
2 | Reconstitute calibrator and place into P01 | 
3 | Place 5 empty tubes into P02 – P06 | 
4 | Wait for 60 mins | 
Procedure |  | 
1 | Open menu SETUP TEST and select FV | 
2 | Enter LOT and expiry date by hand or barcode | 
3 | Select UNIT = % | 
4 | Select ENTRY=AUTO | 
5 | Enter reference value of calibrator | 
6 | Run measurement | 
7 | Verify and save calibration curve | 
Factor V Deficient Plasma
REF: P5101-010

## Slide 26
Instrument | Coatron A6 Plus | 
Firmware | V1.03.05d (50) | 
 | Volume | Position
PAT | 8 | 
BUF | 32 | P79
CLR | 0 | 
DEF | 40 | P62
R0 | 0 | 
R1 | 0 | 
R2 | 80 | P46
REF | Determinations | 
P5201-010 | 10x1ml | n=250
Test | F7
UNIT | 32771
S-CORR | 0%
T-CORR | 0%, 0 s
START | 9 s
INCUB | 120 s
RUNTIME | 180 s
METHOD | COAG
MATH | logXY
CT-MECH | NO
SENS | HI
MIX | 0
CLEAN | 0
POINTS | 6
MULTI | 0
Factor VII Deficient Plasma
REF: P5201-010
Test Protocol
P46: PT reagent
P62: Factor VII deficient plasma
P79: IBS buffer

## Slide 27
Auto Calibration of FVII-% |  |  | 
Material |  |  | 
Calibrator | TECAL N |  | 1ml
Diluent | IBS Buffer |  | 5ml
Tubes | 13x75 mm |  | 6x
Preparation |  |  | 
1 | Reconstitute Factor VII deficient plasma and PT reagent, place them onboard |  | 
2 | Reconstitute calibrator and place into P01 |  | 
3 | Place 5 empty tubes into P02 – P06 |  | 
4 | Wait for 60 mins |  | 
Procedure |  |  | 
1 | Open menu SETUP TEST and select FVII |  | 
2 | Enter LOT and expiry date by hand or barcode |  | 
3 | Select UNIT = % |  | 
4 | Select ENTRY=AUTO |  | 
5 | Enter reference value of calibrator |  | 
6 | Run measurement |  | 
7 | Verify and save calibration curve |  | 
Typical calibration curve (The laboratory must establish its own curve for each new LOT of reagent ) |  |  | 
1 | 108.0% | 19.3 sec | 
2 | 54.0% | 22.5 sec | 
3 | 27.0% | 27.6 sec | 
4 | 13.5% | 33.8 sec | 
5 | 6.8% | 39.6 sec | 
6 | 1.3% | 54.3 sec | 
Factor VII Deficient Plasma
REF: P5201-010

## Slide 28
Instrument | Coatron A6 Plus | 
Firmware | V1.03.05d (50) | 
 | Volume | Position
PAT | 8 | 
BUF | 32 | P79
CLR | 0 | 
DEF | 40 | P63
R0 | 0 | 
R1 | 40 | P60
R2 | 40 | P47
REF | Determinations | 
P5301-010 | 10x1ml | n=250
Test | F8
UNIT | 32771
S-CORR | 0%
T-CORR | 0%, 0 s
START | 21 s
INCUB | 180 s
RUNTIME | 180 s
METHOD | COAG
MATH | logXY
CT-MECH | NO
SENS | LO
MIX | 0
CLEAN | 0
POINTS | 6
MULTI | 0
Factor VIII Deficient Plasma
REF: P5301-010
Test Protocol
P47: CaCl2
P60: aPTT reagent
P63: Factor VIII deficient plasma
P79: IBS buffer

## Slide 29
Auto Calibration of FVIII-% |  |  | 
Material |  |  | 
Calibrator | TECAL N |  | 1ml
Diluent | IBS Buffer |  | 5ml
Tubes | 13x75 mm |  | 6x
Preparation |  |  | 
1 | Reconstitute Factor VIII deficient plasma, place it and aPTT, CaCl2 onboard |  | 
2 | Reconstitute calibrator and place into P01 |  | 
3 | Place 5 empty tubes into P02 – P06 |  | 
4 | Wait for 30 mins |  | 
Procedure |  |  | 
1 | Open menu SETUP TEST and select FVIII |  | 
2 | Enter LOT and expiry date by hand or barcode |  | 
3 | Select UNIT = % |  | 
4 | Select ENTRY=AUTO |  | 
5 | Enter reference value of calibrator |  | 
6 | Run measurement |  | 
7 | Verify and save calibration curve |  | 
Typical calibration curve (The laboratory must establish its own curve for each new LOT of reagent ) |  |  | 
1 | 71.0% | 54.5 sec | 
2 | 35.5% | 60.7 sec | 
3 | 17.7% | 68.1 sec | 
4 | 8.8% | 75.9 sec | 
5 | 4.4% | 85.1 sec | 
6 | 2.2% | 91.9 sec | 
Factor VIII Deficient Plasma
REF: P5301-010

## Slide 30
Instrument | Coatron A6 Plus | 
Firmware | V1.03.05d (50) | 
 | Volume | Position
PAT | 8 | 
BUF | 32 | P79
CLR | 0 | 
DEF | 40 | P66
R0 | 0 | 
R1 | 40 | P60
R2 | 40 | P47
REF | Determinations | 
P5401-010 | 10x1ml | n=250
Test | F9
UNIT | 32771
S-CORR | 0%
T-CORR | 0%, 0 s
START | 21 s
INCUB | 180 s
RUNTIME | 180 s
METHOD | COAG
MATH | logXY
CT-MECH | NO
SENS | LO
MIX | 0
CLEAN | 0
POINTS | 6
MULTI | 0
Factor IX Deficient Plasma
REF: P5401-010
Test Protocol
P47: CaCl2
P60: aPTT reagent
P66: Factor IX deficient plasma
P79: IBS buffer

## Slide 31
Auto Calibration of FIX-% |  | 
Material |  | 
Calibrator | TECAL N | 1ml
Diluent | IBS Buffer | 5ml
Tubes | 13x75 mm | 6x
Preparation |  | 
1 | Reconstitute Factor IX deficient plasma, place it and aPTT, CaCl2 onboard | 
2 | Reconstitute calibrator and place into P01 | 
3 | Place 5 empty tubes into P02 – P06 | 
4 | Wait for 30 mins | 
Procedure |  | 
1 | Open menu SETUP TEST and select FIX | 
2 | Enter LOT and expiry date by hand or barcode | 
3 | Select UNIT = % | 
4 | Select ENTRY=AUTO | 
5 | Enter reference value of calibrator | 
6 | Run measurement | 
7 | Verify and save calibration curve | 
Factor IX Deficient Plasma
REF: P5401-010
Typical calibration curve (The laboratory must establish its own curve for each new LOT of reagent ) |  |  | 
1 | 83.0% | 56.4 sec | 
2 | 41.5% | 60.5 sec | 
3 | 20.7% | 66.9 sec | 
4 | 10.3% | 75.2 sec | 
5 | 5.1% | 85.8 sec | 
6 | 2.5% | 93.1 sec | 

## Slide 32
Instrument | Coatron A6 Plus | 
Firmware | V1.03.05d (50) | 
 | Volume | Position
PAT | 8 | 
BUF | 32 | P79
CLR | 0 | 
DEF | 40 | P62
R0 | 0 | 
R1 | 0 | 
R2 | 80 | P46
REF | Determinations | 
P5501-010 | 10x1ml | n=250
Test | F10
UNIT | 32771
S-CORR | 0%
T-CORR | 0%, 0 s
START | 9 s
INCUB | 120 s
RUNTIME | 180 s
METHOD | COAG
MATH | logXY
CT-MECH | NO
SENS | HI
MIX | 0
CLEAN | 0
POINTS | 6
MULTI | 0
Factor X Deficient Plasma
REF: P5501-010
Test Protocol
P46: PT reagent
P62: Factor X deficient plasma
P79: IBS buffer

## Slide 33
Auto Calibration of FX-% |  | 
Material |  | 
Calibrator | TECAL N | 1ml
Diluent | IBS Buffer | 5ml
Tubes | 13x75 mm | 6x
Preparation |  | 
1 | Reconstitute Factor X deficient plasma and PT reagent, place them onboard | 
2 | Reconstitute calibrator and place into P01 | 
3 | Place 5 empty tubes into P02 – P06 | 
4 | Wait for 30 mins | 
Procedure |  | 
1 | Open menu SETUP TEST and select FX | 
2 | Enter LOT and expiry date by hand or barcode | 
3 | Select UNIT = % | 
4 | Select ENTRY=AUTO | 
5 | Enter reference value of calibrator | 
6 | Run measurement | 
7 | Verify and save calibration curve | 
Factor X Deficient Plasma
REF: P5501-010

## Slide 34
Instrument | Coatron A6 Plus | 
Firmware | V1.03.05d (50) | 
 | Volume | Position
PAT | 8 | 
BUF | 32 | P79
CLR | 0 | 
DEF | 40 | P67
R0 | 0 | 
R1 | 40 | P60
R2 | 40 | P47
REF | Determinations | 
P5601-010 | 10x1ml | n=250
Test | F11
UNIT | 32771
S-CORR | 0%
T-CORR | 0%, 0 s
START | 21 s
INCUB | 180 s
RUNTIME | 180 s
METHOD | COAG
MATH | logXY
CT-MECH | NO
SENS | LO
MIX | 0
CLEAN | 0
POINTS | 6
MULTI | 0
Factor XI Deficient Plasma
REF: P5601-010
Test Protocol
P47: CaCl2
P60: aPTT reagent
P67: Factor XI deficient plasma
P79: IBS buffer

## Slide 35
Auto Calibration of FXI-% |  | 
Material |  | 
Calibrator | TECAL N | 1ml
Diluent | IBS Buffer | 5ml
Tubes | 13x75 mm | 6x
Preparation |  | 
1 | Reconstitute Factor XI deficient plasma, place it and aPTT, CaCl2 onboard | 
2 | Reconstitute calibrator and place into P01 | 
3 | Place 5 empty tubes into P02 – P06 | 
4 | Wait for 30 mins | 
Procedure |  | 
1 | Open menu SETUP TEST and select FXI | 
2 | Enter LOT and expiry date by hand or barcode | 
3 | Select UNIT = % | 
4 | Select ENTRY=AUTO | 
5 | Enter reference value of calibrator | 
6 | Run measurement | 
7 | Verify and save calibration curve | 
Factor XI Deficient Plasma
REF: P5601-010

## Slide 36
Instrument | Coatron A6 Plus | 
Firmware | V1.03.05d (50) | 
 | Volume | Position
PAT | 8 | 
BUF | 32 | P79
CLR | 0 | 
DEF | 40 | P67
R0 | 0 | 
R1 | 40 | P60
R2 | 40 | P47
REF | Determinations | 
P5701-010 | 10x1ml | n=250
Test | F12
UNIT | 32771
S-CORR | 0%
T-CORR | 0%, 0 s
START | 21 s
INCUB | 180 s
RUNTIME | 180 s
METHOD | COAG
MATH | logXY
CT-MECH | NO
SENS | LO
MIX | 0
CLEAN | 0
POINTS | 6
MULTI | 0
Factor XII Deficient Plasma
REF: P5701-010
Test Protocol
P47: CaCl2
P60: aPTT reagent
P67: Factor XII deficient plasma
P79: IBS buffer

## Slide 37
Auto Calibration of FXII-% |  | 
Material |  | 
Calibrator | TECAL N | 1ml
Diluent | IBS Buffer | 5ml
Tubes | 13x75 mm | 6x
Preparation |  | 
1 | Reconstitute Factor XII deficient plasma, place it and aPTT, CaCl2 onboard | 
2 | Reconstitute calibrator and place into P01 | 
3 | Place 5 empty tubes into P02 – P06 | 
4 | Wait for 30 mins | 
Procedure |  | 
1 | Open menu SETUP TEST and select FXII | 
2 | Enter LOT and expiry date by hand or barcode | 
3 | Select UNIT = % | 
4 | Select ENTRY=AUTO | 
5 | Enter reference value of calibrator | 
6 | Run measurement | 
7 | Verify and save calibration curve | 
Factor XII Deficient Plasma
REF: P5701-010

## Slide 38
Instrument | Coatron A6 Plus | 
Firmware | V1.03.05d (50) | 
 | Volume | Position
PAT | 8 | 
BUF | 24 | P79
CLR | 0 | 
DEF | 0 | 
R0 | 0 | 
R1 | 32 | P68
R2 | 32 | P56
REF | Determinations | 
C1100-012 | kit | n=375
Test | PC
UNIT | 32771
S-CORR | 0%
T-CORR | 0%, 0 s
START | 60 s
INCUB | 300 s
RUNTIME | 180 s
METHOD | CHROM
MATH | lin
CT-MECH | NO
LIMIT | 250 mE
MIX | 0
CLEAN | 3
POINTS | 3
MULTI | 0
TEChrom Protein C
REF: C1100-012
Test Protocol
P68: Protein C Activator
P56: Protein C Substrate
P79: IBS buffer

## Slide 39
Auto Calibration of PC-%, IU/mL |  | 
Material |  | 
Calibrator | TECAL N | 1ml
Diluent | IBS Buffer | 5ml
Tubes | 13x75 mm | 3x
Preparation |  | 
1 | Reconstitute Protein C reagents and place them onboard | 
2 | Reconstitute calibrator and place into P01 | 
3 | Place 2 empty tubes into P02 and P03 | 
4 | Wait for 30 mins | 
Procedure |  | 
1 | Open menu SETUP TEST and select PC | 
2 | Enter LOT and expiry date by hand or barcode | 
3 | Select UNIT = % or IU/mL | 
4 | Select ENTRY=AUTO | 
5 | Enter reference value of calibrator | 
6 | Run measurement | 
7 | Verify and save calibration curve | 
TEChrom Protein C
REF: C1100-012
Typical calibration curve (The laboratory must establish its own curve for each new LOT of reagent ) |  | 
1 | 104.0% | 0.072 E
2 | 52.0% | 0.039 E
3 | 26.0% | 0.021 E

## Slide 40
Instrument | Coatron A6 Plus | 
Firmware | V1.03.05d (50) | 
 | Volume | Position
PAT | 5 | 
BUF | 35 | P76
CLR | 0 | 
DEF | 35 | P63
R0 | 35 | P66
R1 | 35 | P54
R2 | 35 | P47
REF | Determinations | 
A0600-002 | kit | n=57
Test | PS
UNIT | 32771
S-CORR | 0%
T-CORR | 0%, 0 s
START | 21 s
INCUB | 180 s
RUNTIME | 450 s
METHOD | COAG
MATH | lin
CT-MECH | NO
SENS | HI
MIX | 0
CLEAN | 3
POINTS | 3
MULTI | 0
TECLOT Protein S
REF: A0600-002
Test Protocol
P76: 0.9% Saline solution
P63: Protein S Deficient plasma
P66: Protein S Activator
P54: Protein S Substrate
P47: 25mM CaCl2

## Slide 41
Auto Calibration of PS-%, IU/mL |  | 
Material |  | 
Calibrator | TECAL N | 1ml
Diluent | IBS Buffer | 5ml
Tubes | 13x75 mm | 3x
Preparation |  | 
1 | Reconstitute Protein S reagents and place them onboard | 
2 | Reconstitute calibrator and place into P01 | 
3 | Place 2 empty tubes into P02 and P03 | 
4 | Wait for 30 mins | 
Procedure |  | 
1 | Open menu SETUP TEST and select PS | 
2 | Enter LOT and expiry date by hand or barcode | 
3 | Select UNIT = % or IU/mL | 
4 | Select ENTRY=AUTO | 
5 | Enter reference value of calibrator | 
6 | Run measurement | 
7 | Verify and save calibration curve | 
TECLOT Protein S
REF: A0600-002
Typical calibration curve (The laboratory must establish its own curve for each new LOT of reagent ) |  | 
1 | 96.0% | 182.5 sec
2 | 48.0% | 163.5 sec
3 | 24.0% | 137.8 sec

## Slide 42
Instrument | Coatron A6 Plus |  | 
Firmware | V1.03.05d (50) |  | 
 | Volume | Position | 
PAT | 10 |  | 
BUF | 0 | P79 | 
CLR | 0 |  | 
DEF | 40 | P59 | 
R0 | 0 |  | 
R1 | 50 | P60 (APTT) | 
 | 50 | P55 (PCA.APTT) | 
R2 | 50 | P47 | 
REF | Determinations |  | 
A0900-004 | kit |  | n=80
Test | - APC | APCR
UNIT | 1 | 1
S-CORR | 0% | 0%
T-CORR | 0%, 0 s | 0%, 0 s
START | 21 s | 21 s
INCUB | 300 s | 300 s
RUNTIME | 300 s | 300 s
METHOD | COAG | COAG
MATH | lin | lin
CT-MECH | NO | NO
SENS | LO | LO
MIX | 0 | 0
CLEAN | 0 | 3
POINTS | 0 | 0
MULTI | 0 | 0
TECLOT PCA Ratio
REF: A0900-004
Test Protocol
P47: CaCl2
P60: aPTT reagent
P59: Factor V Depleted Plasma
P55: PCA.APTT reagent

## Slide 43
PCA- Ratio | 
Preparation | 
1 | Reconstitute PCA reagents and place them onboard
2 | Wait for 30 mins
Procedure | 
1 | Open menu SETUP TEST and select APCA
2 | Enter LOT and expiry date by hand or barcode
3 | Run measurement
4 | Result in Ratio will be automatically calculated by instrument: R=Clot time (APCR)/Clot time (-APC)
TECLOT PCA Ratio
REF: A0900-004
Interpretation R=Clot time (APCR)/Clot time (-APC) | 
Normality | R > 2.2
FV leiden mutation | R < 2.2

## Slide 44
Instrument | Coatron A6 Plus | 
Firmware | V1.03.05d (50) | 
 | Volume | Position
PAT | 38 | 
BUF | 0 | P79
CLR | 0 | 
DEF | 0 | 
R0 | 0 | 
R1 | 0 | 
R2 | 38 | P51
REF | Determinations | 
A0700-020 | 10x2ml | n=526
Test | LA-S
UNIT | 17
S-CORR | 0%
T-CORR | 0%, 0 s
START | 11 s
INCUB | 60 s
RUNTIME | 240 s
METHOD | COAG
MATH | lin
CT-MECH | NO
SENS | LO
MIX | 0
CLEAN | 3
POINTS | 0
MULTI | 0
TECLOT LA Screen
REF: A0700-020
Test Protocol
P51: LA-Screen reagent

## Slide 45
LA-S- Ratio | 
Preparation | 
1 | Reconstitute LA-S reagent and place it onboard
2 | Wait for 30 mins
Procedure | 
1 | Open menu SETUP TEST and select LA-S
2 | Enter LOT and expiry date by hand or barcode
3 | Select UNIT = R
4 | Select ENTRY=Manual
5 | Enter Normal value
6 | Verify and save
TECLOT LA Screen
REF: A0700-020
Interpretation If LA-S Ratio is greater than 1.2, a lupus anticoagulant may be present. Confirm lupus anticoagulant with test LA-C (Cat. No. A0800-020) : Lupus Ratio (LR) = Ratio of LA-S / Ratio of LA-C | 
Strong LA | LR > 2.0
Moderate LA | LR = 1.5 -2.0
Weak LA | LR = 1.2 -1.5

## Slide 46
Instrument | Coatron A6 Plus | 
Firmware | V1.03.05d (50) | 
 | Volume | Position
PAT | 38 | 
BUF | 0 | P79
CLR | 0 | 
DEF | 0 | 
R0 | 0 | 
R1 | 0 | 
R2 | 38 | P48
REF | Determinations | 
A0700-020 | 10x1ml | n=263
Test | LA-C
UNIT | 17
S-CORR | 0%
T-CORR | 0%, 0 s
START | 11 s
INCUB | 60 s
RUNTIME | 240 s
METHOD | COAG
MATH | lin
CT-MECH | NO
SENS | LO
MIX | 0
CLEAN | 3
POINTS | 0
MULTI | 0
TECLOT LA Confirm
REF: A0800-010
Test Protocol
P48: LA Confirm reagent

## Slide 47
LA-C- Ratio | 
Preparation | 
1 | Reconstitute LA-C reagent and place it onboard
2 | Wait for 30 mins
Procedure | 
1 | Open menu SETUP TEST and select LA-C
2 | Enter LOT and expiry date by hand or barcode
3 | Select UNIT = R
4 | Select ENTRY=Manual
5 | Enter Normal value
6 | Verify and save
Interpretation If LA-S Ratio is greater than 1.2, a lupus anticoagulant may be present. Confirm lupus anticoagulant with test LA-C (Cat. No. A0800-020) : Lupus Ratio (LR) = Ratio of LA-S / Ratio of LA-C | 
Strong LA | LR > 2.0
Moderate LA | LR = 1.5 -2.0
Weak LA | LR = 1.2 -1.5
TECLOT LA Confirm
REF: A0800-010

## Slide 48: Rückblatt Seite-1

## Slide 49: Rückblatt-2
 | Copyright © 2018, TECO Medical Instruments, Production + Trading GmbH Dieselstrasse. 1; D-84088 Neufahrn in N.B in Germany