## Slide 1
Instrument | Coatron A4 | 
Firmware | V1.07.05a (261) | 
 | Volume | Position
PAT | 5 | 35
BUF | 45 | 39
CLR | 0 | 
DEF | 0 | 
R0 | 0 | 
R1 | 49 | 32
R2 | 49 | 28
REF | Determinations | 
 | kit | n=80
Test | AF10
UNIT | 65283
S-CORR | 0%
T-CORR | 0%, 0 s
START | 20 s
INCUB | 60 s
RUNTIME | 60 s
METHOD | CHROM
MATH | LIN
CT-MECH | NO
LIMIT | 350mE
MIX | NO
CLEAN | 1
POINTS | 4
MULTI | NO
Technochrom anti-Xa
for LMWH, UHF, Arixatra
Test Protocol HEP
P32: Substrate (R1)
P28: Factor Xa (R2)
P39: IBS

## Slide 2
Auto Calibration |  | 
Material |  | 
Calibrator | TECHNOVIEW LMWH Con-H | 1ml
Diluent | TEControl N | 1ml
Tubes | 13x75 mm | 4x
Preparation |  | 
1 | Place P01: 400µL LMWH Con-H | 
2 | Place P02: 200µL LMWH Con-H + 200µl TEControl-N | 
3 | Place P03: 100µL LMWH Con-H + 300µl TEControl-N | 
4 | Place P04: 400µL TEControl-N | 
Procedure |  | 
1 | Open menu SETUP TEST and select HEP | 
2 | Enter LOT and expiry date by hand or barcode | 
3 | Select UNIT = IE/ml | 
4 | Select ENTRY=AUTO | 
5 | Enter STD1 = 1.04 IE/mL (value of LMWH Con-H - example) | 
6 | Enter STD2 = 0.52 IE/mL (50% value of LMWH Con-H) | 
7 | Enter STD3 = 0.26 IE/mL (25% value of LMWH Con-H) | 
8 | Enter STD4 = 0.01 IE/mL ( value of TEControl-N) | 
9 | Run measurement | 
10 | Verify and save calibration curve | 
TEChrom AF10
for LMWH, UHF, Arixatra

## Slide 3
Instrument | Coatron A4 | 
Firmware | V1.07.05a (261) | 
 | Volume | Position
PAT | 7 | 35
BUF | 42 | 39
CLR | 0 | 
DEF | 0 | 
R0 | 0 | 
R1 | 49 | 32
R2 | 49 | 28
REF | Determinations | 
 | kit | n=80
Test | AF10
UNIT | 65283
S-CORR | 0%
T-CORR | 0%, 0 s
START | 20 s
INCUB | 60 s
RUNTIME | 60 s
METHOD | CHROM
MATH | LIN
CT-MECH | NO
LIMIT | 350mE
MIX | NO
CLEAN | 1
POINTS | 4
MULTI | NO
Technochrom anti-Xa
for Edoxaban
Test Protocol AF10
P32: Substrate (R1)
P28: Factor Xa (R2)
P39: IBS

## Slide 4
Auto Calibration |  | 
Material |  | 
Calibrator | TECHNOVIEW EDOXABAN Con-H | 1ml
Diluent | TEControl N | 1ml
Tubes | 13x75 mm | 4x
Preparation |  | 
1 | Place P01: 400µL EDOXABAN Con-H | 
2 | Place P02: 200µL EDOXABAN Con-H + 200µl TEControl-N | 
3 | Place P03: 100µL EDOXABAN Con-H + 300µl TEControl-N | 
4 | Place P04: 400µL TEControl-N | 
Procedure |  | 
1 | Open menu SETUP TEST and select HEP | 
2 | Enter LOT and expiry date by hand or barcode | 
3 | Select UNIT = ng/ml | 
4 | Select ENTRY=AUTO | 
5 | Enter STD1 = 324 ng/ml (value of EDOXABAN Con-H - example) | 
6 | Enter STD2 = 171 ng/ml (50% value of EDOXABAN Con-H) | 
7 | Enter STD3 = 86 ng/ml (25% value of EDOXABAN Con-H) | 
8 | Enter STD4 = 1 ng/ml ( value of TEControl-N) | 
9 | Run measurement | 
10 | Verify and save calibration curve | 
TEChrom AF10
for LMWH, UHF, Arixatra